EPID Gain Map Pipeline

Building a Detector Sensitivity Correction for MPC Image Analysis

Clinical Physics  •  MPC Capstone Project

# 1. Background: Why You Need a Gain Map

The Varian AS1200 electronic portal imaging device (EPID) is an amorphous silicon (aSi) flat-panel detector. Like all detectors, individual pixels have slightly different sensitivities. The raw image you acquire contains two things mixed together: the actual radiation beam profile (the signal you want to measure) and the detector’s pixel-to-pixel sensitivity variations (noise you want to remove).

The purpose of a gain map (also called a pixel sensitivity map) is to isolate and correct the detector’s hardware artifacts without destroying the beam characteristics you are trying to measure, such as flatness and symmetry.

## 1.1 Why Not Just Divide by the Flood Field Directly?

A common first instinct is to normalize a clinical image by dividing it by a flood field image. The problem is that MPC acquires its flood field using a large open field (e.g., 40×40 cm²) delivered through the standard flattening filter. This field is not uniform. It contains the beam’s inherent shape: the characteristic “horns” near the field edges, off-axis softening, and any real asymmetry in the beam. If you divide your clinical image by this unfiltered flood, you remove both the detector artifacts and the beam shape—flattening out the very signal you are trying to measure.

This is well-documented in EPID dosimetry literature: standard flood-field correction suppresses beam horns and profile variations because those low-frequency features exist in both the flood and the clinical image, and they cancel out in the division.

# 2. Core Concept: Frequency Separation

The solution is to separate the flood field image into two components based on spatial frequency:

- Low-frequency component: The broad beam shape (horns, flatness profile, any real asymmetry). These features vary slowly over hundreds of pixels.

- High-frequency component: The detector’s pixel-to-pixel sensitivity variations, dead pixels, and readout line artifacts. These are sharp, localized features that change over just a few pixels.

By applying a large-kernel smoothing filter to the flood field, you produce an estimate of just the beam shape. Dividing the original flood by this smoothed version isolates only the high-frequency detector artifacts. This resulting image is your gain map: it has a mean value near 1.0 and contains only the “speckle” and line patterns of the detector hardware.

When you then divide a clinical image by this gain map, you correct the detector artifacts while leaving the clinical beam’s shape completely intact.

# 3. The Pipeline, Step by Step

The complete processing pipeline has two phases: an offline phase (building the master gain map, done periodically) and a daily phase (correcting each day’s clinical image for analysis).

## 3.1 Offline Phase: Building the Master Gain Map

- Acquire Dark Frame(s). A dark frame is an image acquired with no radiation beam. It captures the detector’s electronic baseline offset. Subtract this from every subsequent image to remove the electronic pedestal. If your MPC data includes a dark frame for each session, use it; otherwise, use the most recent available dark frame.

- Collect Multiple Flood Fields. Gather the last 5–10 flood field images from your MPC acquisitions. Using multiple floods is critical because any single flood contains random noise (photon statistics, transient artifacts, cosmic ray hits). A single flood’s noise would be “baked into” your gain map and contaminate every clinical image you correct with it.

- Dark-Subtract Each Flood. For each flood image, subtract the corresponding dark frame:  F_net[i] = F_raw[i] − Dark[i]. This gives you a set of net flood images that represent the detector’s response to the beam, free of electronic offset.

- Median-Stack the Net Floods. Combine the net floods into a single master flood by taking the pixel-wise median (not the mean). The median is critical here because it naturally rejects outliers: a cosmic ray hit or transient artifact that appears in one flood but not others will be excluded automatically. The result is a clean, high signal-to-noise master flood:  F_master = median(F_net[1], F_net[2], ..., F_net[n]).

- Estimate the Beam Shape (Low-Pass Filter). Apply a large-kernel median filter to the master flood. This smooths out the pixel-level detector artifacts while preserving the broad beam profile:  B = median_filter(F_master, kernel_size=K). Use a median filter here (not Gaussian) because it is robust to outlier pixels—a Gaussian would smear hot and dead pixel values into the beam shape estimate, subtly biasing your gain map near those defective pixels.

Choosing kernel size K: The kernel must be much larger than the scale of detector artifacts (individual pixels, readout lines ≈ 5–10 pixels) but smaller than the beam profile curvature scale (≈ 200+ pixels). A starting point of K = 75–101 (must be odd) is reasonable for the AS1200’s 0.336 mm pixel pitch. Validate empirically: plot a cross-profile of the resulting gain map. It should be flat (~1.0) in smooth detector regions while preserving the readout line artifacts. If you see residual beam curvature, increase K; if the line artifacts start to blur, decrease K.

- Compute the Gain Map. Divide the master flood by the smoothed beam estimate:  G = F_master / B. The result is your gain map. It should have a mean value very close to 1.0 across the in-field region and contain only high-frequency detector artifacts.

- Normalize Over the In-Field ROI. Normalize the gain map so that its mean equals exactly 1.0, computed only within a robust region of interest (e.g., the central 80% of the field width). Do not use the entire image array for normalization—edge and penumbra pixels will bias the mean away from 1.0 in the clinically relevant region.

- Flag and Clip Dead/Hot Pixels. Pixels in the gain map whose values deviate significantly from 1.0 (e.g., outside the range 0.7–1.3) are dead or hot pixels. Flag these pixels in a boolean mask for later exclusion, and clip the gain map values to your threshold bounds. This prevents divide-by-near-zero blowups when you apply the correction. These flagged pixels should be excluded from any downstream flatness or symmetry calculations.

## 3.2 Daily Phase: Correcting a Clinical Image

- Dark-Subtract the Clinical Image. Subtract the day’s dark frame from the raw clinical acquisition:  C_net = C_raw − Dark_daily.

- Apply the Gain Map Correction. Divide the net clinical image by the master gain map:  C_corrected = C_net / G. This removes the detector’s pixel sensitivity variations while preserving the beam’s true shape.

- Smooth the Beam Profiles Before Metric Extraction. Even after gain correction, the extracted crossline and inline profiles will contain residual pixel-level noise from photon statistics, interpolated bad pixels, and imperfect gain correction. If you calculate flatness and symmetry directly on these raw profiles, the noise inflates your max and min values, making your metrics unstable from day to day. Before computing any beam metrics, apply a light smoothing filter to the 1D extracted profiles. A small-kernel Savitzky–Golay filter (e.g., window length 11–21, polynomial order 2–3) is a good choice because it preserves the shape of peaks and shoulders while suppressing point noise. Alternatively, a simple moving average or small-kernel median on the 1D profile works. The key is to keep the smoothing kernel small relative to the beam profile features—you are removing pixel noise, not reshaping the beam.

Important: This smoothing is applied to the extracted 1D profiles, not to the 2D image. You want the full-resolution image for spatial analysis and trending, but the 1D profiles used for flatness and symmetry calculations benefit from noise suppression to produce stable, reproducible metrics.

- Analyze Within the In-Field ROI. Perform your beam analysis (flatness, symmetry, etc.) on the smoothed profiles, restricting measurements to the in-field region (typically within 80% of the field width) and excluding any pixels that were flagged as dead or hot in the gain map. Penumbra regions should be avoided unless you are specifically studying penumbra behavior.

# 4. Pipeline Summary

| Step | Action | Detail | Result |
| --- | --- | --- | --- |
| 1 | Acquire Dark | Image with no beam (electronic baseline) | Dark frame |
| 2 | Collect Floods | 5–10 recent flood field images | Flood set |
| 3 | Dark-Subtract Floods | F_net[i] = F_raw[i] − Dark[i] | Net floods |
| 4 | Median-Stack | Pixel-wise median across all net floods | F_master |
| 5 | Low-Pass Filter | Large-kernel median filter on F_master | Beam shape (B) |
| 6 | Compute Gain Map | G = F_master / B | Gain map (G) |
| 7 | Normalize G | Mean = 1.0 over in-field ROI | Normalized G |
| 8 | Clip Extremes | Flag pixels outside 0.7–1.3; build mask | Clean G + mask |
| 9 | Dark-Subtract Clinical | C_net = C_raw − Dark_daily | Net clinical |
| 10 | Apply Correction | C_corrected = C_net / G | Corrected image |
| 11 | Extract & Smooth Profiles | Savitzky–Golay or moving average on 1D profiles | Clean profiles |
| 12 | Analyze | Flatness, symmetry within 80% field | QA metrics |

# 5. How to Validate Your Pipeline

Before trusting your pipeline for clinical constancy tracking, run these sanity checks:

- Profile Preservation Test: Plot crossline and inline profiles of the clinical image before and after gain correction. You should see detector line artifacts disappear while the broad beam shape (horns, slope) remains unchanged.

- Residual Artifact Map: Examine the corrected image for any remaining row or column fingerprints. If you see them, your gain map may have insufficient SNR (use more floods) or your kernel size needs adjustment.

- Gain Map Inspection: The gain map itself should look like a “flat” image (mean ~1.0) with only fine-grained detector structure visible. If you see residual beam curvature in the gain map, your filter kernel is too small.

- Dead Pixel Check: Verify that your dead/hot pixel mask captures known bad pixels (you can compare against Varian’s service logs if available). Ensure these pixels are excluded from flatness and symmetry calculations.

# 6. Reference Implementation (Python / pylinac)

The following pseudocode demonstrates the complete pipeline using Python with NumPy, SciPy, and pylinac. This is meant as a structural guide—you will need to adapt file paths, image loading, and pylinac API calls to match your specific data format and pylinac version.

## 6.1 Building the Master Gain Map

import numpy as np

from scipy.ndimage import median_filter

from pathlib import Path

def load_epid_image(filepath):

"""

Load a raw EPID image from an MPC acquisition file.

Adapt this to your file format (DICOM, HIS, etc.).

Returns a 2D numpy array of pixel values.

"""

# Example for DICOM:

# import pydicom

# ds = pydicom.dcmread(filepath)

# return ds.pixel_array.astype(np.float64)

pass

def build_master_gain_map(

flood_paths: list[Path],

dark_path: Path,

kernel_size: int = 75,

clip_low: float = 0.7,

clip_high: float = 1.3,

field_fraction: float = 0.8

) -> tuple[np.ndarray, np.ndarray]:

"""

Build a master gain map from multiple flood field images.

Parameters

----------

flood_paths : list of Paths to flood field images

dark_path   : Path to the dark frame image

kernel_size : Median filter kernel (must be odd)

clip_low    : Lower bound for gain clipping

clip_high   : Upper bound for gain clipping

field_fraction : Fraction of field for ROI normalization

Returns

-------

gain_map    : 2D array, mean ~1.0, detector artifacts only

bad_pixel_mask : Boolean array, True = bad pixel

"""

# --- Step 1: Load the dark frame ---

dark = load_epid_image(dark_path)

# --- Steps 2-3: Dark-subtract each flood ---

floods_net = []

for fpath in flood_paths:

flood_raw = load_epid_image(fpath)

flood_net = flood_raw - dark

floods_net.append(flood_net)

# --- Step 4: Median-stack to reject outliers ---

flood_stack = np.stack(floods_net, axis=0)

f_master = np.median(flood_stack, axis=0)

# --- Step 5: Estimate beam shape via large-kernel median ---

beam_shape = median_filter(f_master, size=kernel_size)

# --- Step 6: Isolate detector sensitivity ---

# Avoid divide-by-zero in low-signal regions

beam_shape_safe = np.where(

beam_shape > 0.01 * beam_shape.max(),

beam_shape,

1.0

)

gain_map = f_master / beam_shape_safe

# --- Step 7: Normalize over in-field ROI ---

rows, cols = gain_map.shape

margin = int((1 - field_fraction) / 2 * min(rows, cols))

roi = gain_map[margin:rows-margin, margin:cols-margin]

gain_map = gain_map / np.mean(roi)

# --- Step 8: Flag and clip dead/hot pixels ---

bad_pixel_mask = (gain_map < clip_low) | (gain_map > clip_high)

gain_map = np.clip(gain_map, clip_low, clip_high)

return gain_map, bad_pixel_mask

## 6.2 Daily Correction and Analysis

from pylinac.core.image import DicomImage

from pylinac import FieldAnalysis

def correct_clinical_image(

clinical_path: Path,

dark_path: Path,

gain_map: np.ndarray,

bad_pixel_mask: np.ndarray

) -> np.ndarray:

"""

Apply gain map correction to a clinical EPID image.

Parameters

----------

clinical_path  : Path to the raw clinical image

dark_path      : Path to the day's dark frame

gain_map       : Master gain map from build_master_gain_map()

bad_pixel_mask : Boolean mask of dead/hot pixels

Returns

-------

corrected : 2D array ready for beam analysis

"""

# --- Step 9: Dark-subtract the clinical image ---

clinical_raw = load_epid_image(clinical_path)

dark = load_epid_image(dark_path)

c_net = clinical_raw - dark

# --- Step 10: Apply gain correction ---

corrected = c_net / gain_map

# --- Optional: Interpolate bad pixels ---

# Replace flagged pixels with local median of neighbors

# so they don't contaminate analysis

if np.any(bad_pixel_mask):

from scipy.ndimage import generic_filter

local_med = generic_filter(

corrected, np.nanmedian, size=5

)

corrected[bad_pixel_mask] = local_med[bad_pixel_mask]

return corrected

## 6.3 Profile Smoothing and Beam Analysis with pylinac

from scipy.signal import savgol_filter

from pylinac.core.image import DicomImage

from pylinac import FieldAnalysis

def smooth_profile(

profile: np.ndarray,

window_length: int = 15,

polyorder: int = 3

) -> np.ndarray:

"""

Smooth a 1D beam profile to suppress pixel-level noise

before calculating flatness and symmetry metrics.

Uses Savitzky-Golay filtering, which preserves peak shape

and shoulders while removing point noise.

Parameters

----------

profile       : 1D array (crossline or inline profile)

window_length : Must be odd. 11-21 is a good range.

Smaller = less smoothing, more noise.

Larger = more smoothing, risk of shape loss.

polyorder     : Polynomial order (2 or 3 typical).

Must be < window_length.

Returns

-------

smoothed : 1D array, same length as input

"""

return savgol_filter(profile, window_length, polyorder)

def extract_and_smooth_profiles(

corrected_image: np.ndarray,

field_fraction: float = 0.8

) -> dict:

"""

Extract crossline and inline profiles from the corrected

image center, then smooth them for metric extraction.

"""

rows, cols = corrected_image.shape

center_row = rows // 2

center_col = cols // 2

# Extract 1D profiles through image center

crossline_raw = corrected_image[center_row, :]

inline_raw    = corrected_image[:, center_col]

# --- Step 11: Smooth profiles before metric extraction ---

crossline = smooth_profile(crossline_raw)

inline    = smooth_profile(inline_raw)

return {

'crossline': crossline,

'crossline_raw': crossline_raw,  # Keep raw for comparison

'inline': inline,

'inline_raw': inline_raw,

}

def analyze_beam(corrected_image: np.ndarray) -> dict:

"""

Run flatness and symmetry analysis on the corrected image

using pylinac's FieldAnalysis.

Adapt this to your pylinac version and preferred

analysis protocol (Varian, IEC, etc.).

"""

# --- Step 12: Beam analysis ---

# Option A: If pylinac can load from a numpy array

# (check your pylinac version's API for this)

#

# Option B: Save corrected image to a temporary

# DICOM file, then load with pylinac

import tempfile

import pydicom

# Create a minimal DICOM with the corrected pixel data

# (You may need to copy metadata from the original file)

tmp = tempfile.NamedTemporaryFile(suffix='.dcm', delete=False)

# ... write corrected_image into DICOM pixel_array ...

# ... save to tmp.name ...

# Load and analyze with pylinac

# NOTE: pylinac may apply its own profile smoothing

# internally. Check your version's documentation.

# If so, you may rely on pylinac's smoothing instead

# of the manual approach above. Either way, verify

# that profiles are smooth before metrics are computed.

fa = FieldAnalysis(tmp.name)

fa.analyze(

# Configure for your protocol:

# - centering method

# - normalization method

# - field width (e.g., in_field_ratio=0.8)

# - analysis type (Varian, IEC, etc.)

)

results = fa.results_data()

return {

'flatness_horizontal': results.flatness_horizontal,

'flatness_vertical':   results.flatness_vertical,

'symmetry_horizontal': results.symmetry_horizontal,

'symmetry_vertical':   results.symmetry_vertical,

}

## 6.4 Putting It All Together

# =============================================

# EXAMPLE: Full workflow from start to finish

# =============================================

from pathlib import Path

# --- Configure paths ---

flood_dir   = Path('path/to/mpc/floods/')

dark_file   = Path('path/to/dark_frame.dcm')

clinical_file = Path('path/to/clinical_image.dcm')

# --- Phase 1: Build the gain map (offline, periodic) ---

flood_files = sorted(flood_dir.glob('*.dcm'))[-7:]  # Last 7 floods

gain_map, bad_mask = build_master_gain_map(

flood_paths=flood_files,

dark_path=dark_file,

kernel_size=75,    # Validate this empirically!

clip_low=0.7,

clip_high=1.3,

field_fraction=0.8

)

# --- Phase 2: Daily correction and analysis ---

corrected = correct_clinical_image(

clinical_path=clinical_file,

dark_path=dark_file,

gain_map=gain_map,

bad_pixel_mask=bad_mask

)

# --- Phase 2b: Extract and smooth profiles ---

profiles = extract_and_smooth_profiles(corrected)

# Optional: plot raw vs smoothed to verify smoothing

# import matplotlib.pyplot as plt

# plt.plot(profiles['crossline_raw'], alpha=0.4, label='Raw')

# plt.plot(profiles['crossline'], label='Smoothed')

# plt.legend(); plt.title('Crossline Profile'); plt.show()

# --- Phase 3: Beam analysis ---

metrics = analyze_beam(corrected)

print(f"Flatness (H): {metrics['flatness_horizontal']:.2f}%")

print(f"Flatness (V): {metrics['flatness_vertical']:.2f}%")

print(f"Symmetry (H): {metrics['symmetry_horizontal']:.2f}%")

print(f"Symmetry (V): {metrics['symmetry_vertical']:.2f}%")

# 7. Common Pitfalls to Avoid

- Using mean instead of median for stacking. The mean will average in outlier artifacts (cosmic rays, transient dead pixels), contaminating your gain map. Always use the pixel-wise median.

- Using a Gaussian filter instead of median for beam shape estimation. A Gaussian smears hot/dead pixel values into the smooth beam estimate, biasing the gain map near defective pixels. The median filter is naturally robust to these point outliers.

- Forgetting to dark-subtract before stacking. If you median-stack raw floods without dark subtraction, the electronic offset (which can drift day to day) adds noise to your gain map.

- Normalizing over the entire image. Edge and penumbra pixels will pull the gain map’s mean away from 1.0 in the region that matters. Always normalize over the in-field ROI.

- Not clipping extreme gain values before applying the correction. Dividing by a near-zero gain value (dead pixel) produces enormous spikes. Clip first, then apply.

- Using a single flood field. One flood’s random photon noise gets permanently embedded in every corrected image. Use at least 5 floods for your median stack.

- Analyzing into the penumbra region. Penumbra has steep dose gradients where even small alignment uncertainties cause large metric swings. Restrict analysis to within 80% of field width unless you have a specific reason to include penumbra.

- Skipping profile smoothing before metric extraction. Flatness and symmetry formulas use max and min values along a profile. Even a single noisy pixel can shift your max or min and produce unstable day-to-day metrics. A light Savitzky–Golay or moving-average filter on the 1D profiles eliminates this without altering the beam shape.

# 8. The Key Takeaway

The entire pipeline rests on one principle: separate what the detector does from what the beam does. The flood field contains both. A large-kernel median filter separates them by spatial frequency. The gain map captures only the detector; dividing by it corrects the detector without disturbing the beam. A light profile smoothing step then ensures your flatness and symmetry metrics are stable rather than noise-driven. Everything else—stacking, clipping, ROI normalization—is engineering to make those ideas robust in practice.