EPID Gain Map Pipeline

Quick Reference Guide

See the full pipeline document for detailed explanations, validation tests, and pseudocode.

# The Problem

MPC flood fields contain both the beam’s real shape (horns, flatness profile) and the detector’s pixel-to-pixel sensitivity variations. Dividing a clinical image by the raw flood removes both—destroying the beam signal you are trying to measure. The solution is to build a gain map that contains only the detector artifacts, so the beam shape is preserved after correction.

# The Approach: Frequency Separation

Detector artifacts are high-frequency (pixel-scale). The beam shape is low-frequency (hundreds of pixels). A large-kernel median filter on the flood field estimates the beam shape. Dividing the flood by that estimate isolates the detector-only component. That’s your gain map.

# Pipeline at a Glance

| # | Step | What to Do | Output |
| --- | --- | --- | --- |
| 1 | Dark-Subtract Floods | F_net[i] = F_raw[i] − Dark | Net floods |
| 2 | Median-Stack | Pixel-wise median across 5–10 net floods | F_master |
| 3 | Estimate Beam Shape | Large-kernel median filter (K ≈ 75–101, odd) | B |
| 4 | Compute Gain Map | G = F_master / B | G (≈ 1.0 mean) |
| 5 | Normalize & Clip | Mean = 1.0 over 80% field ROI; clip 0.7–1.3 | Clean G + mask |
| 6 | Dark-Subtract Clinical | C_net = C_raw − Dark_daily | Net clinical |
| 7 | Apply Gain Correction | C_corrected = C_net / G | Corrected image |
| 8 | Smooth Profiles | Savitzky–Golay on extracted 1D profiles | Clean profiles |
| 9 | Analyze | Flatness & symmetry within 80% field width | QA metrics |

# Key Implementation Decisions

## Why median (not mean) for stacking?

Median rejects outliers (cosmic rays, transient artifacts). Mean averages them in, contaminating your gain map.

## Why median filter (not Gaussian) for beam shape?

Gaussian smears hot/dead pixel values into the beam estimate, biasing the gain map near defective pixels. Median is robust to point outliers—exactly what you are trying to isolate.

## Why normalize over the in-field ROI only?

Edge and penumbra pixels pull the gain map’s mean away from 1.0 in the region that matters for flatness and symmetry.

## Why clip extreme gain values before applying?

Dead pixels produce near-zero gain values. Dividing by near-zero creates enormous spikes. Clip first, then flag those pixels for exclusion from analysis.

## Why smooth profiles before computing metrics?

Flatness and symmetry use max/min values along a profile. Residual pixel noise inflates these extremes, producing unstable day-to-day metrics. A light Savitzky–Golay filter (window 11–21, order 2–3) on the 1D profiles removes noise without reshaping the beam. Smooth the extracted 1D profiles, not the 2D image.

# Condensed Pseudocode

import numpy as np

from scipy.ndimage import median_filter

from scipy.signal import savgol_filter

# === BUILD GAIN MAP (offline, periodic) ===

dark = load_image(dark_path)

floods = [load_image(f) - dark for f in flood_paths[-7:]]

f_master = np.median(np.stack(floods), axis=0)

beam_shape = median_filter(f_master, size=75)       # tune K empirically

gain_map = f_master / beam_shape

# Normalize over in-field ROI

m = int(0.1 * min(gain_map.shape))                  # 80% field margin

roi = gain_map[m:-m, m:-m]

gain_map /= np.mean(roi)

# Flag and clip dead/hot pixels

bad_mask = (gain_map < 0.7) | (gain_map > 1.3)

gain_map = np.clip(gain_map, 0.7, 1.3)

# === DAILY CORRECTION ===

c_net = load_image(clinical_path) - load_image(dark_daily_path)

c_corrected = c_net / gain_map

# === SMOOTH PROFILES & ANALYZE ===

crossline = c_corrected[c_corrected.shape[0]//2, :]

inline    = c_corrected[:, c_corrected.shape[1]//2]

crossline_smooth = savgol_filter(crossline, 15, 3)

inline_smooth    = savgol_filter(inline, 15, 3)

# Compute metrics within 80% field width on smoothed profiles

# Use pylinac FieldAnalysis or manual calculation

# Pre-Analysis Checklist

- Dark frame acquired and subtracted from all images

- At least 5 flood fields in the median stack

- Median filter kernel validated (gain map flat in smooth regions, line artifacts preserved)

- Gain map normalized over in-field ROI (mean = 1.0)

- Dead/hot pixels clipped and masked

- Profile smoothing applied to 1D profiles before metric extraction

- Analysis restricted to within 80% field width

- Profile preservation test passed (beam shape unchanged after correction)